## _Read me
Direct measurement at λmax
Family: Zero-order spectra (absorbance)
The idea in plain words
If only one compound absorbs light at a certain wavelength, the absorbance there tells you its concentration directly — exactly as for a pure solution.
What is in this workbook
• Sheet 'Standards': 5 spectra (wavelength in column A, one spectrum per column; the column title holds the concentrations)
• Sheet 'Mixtures': 5 spectra (wavelength in column A, one spectrum per column; the column title holds the concentrations)
• Sheet '_Calculation': the calibration and the results done with Excel formulas, so every number can be followed.
• Sheet '_Expected results': what Spectro gives for this example.
• Sheets whose name starts with _ are notes; Spectro does not import them.
How to use it in Spectro
1. File → New project…
2. Edit → Compounds… and add: A, B
3. File → Import spectra… → Add files… → choose this workbook.
4. Press 'Fill concentrations from names'. The Role column is filled from the sheet names (Standards → standard, Mixtures → mixture, …). Press Import.
5. Methods → Univariate calibration → template Direct (zero order, λmax).
6. Choose the compound, then drag the λ marker to a wavelength where only it absorbs (check the pure spectra of the other compounds).
7. Check the calibration standards and the mixtures → Calibrate → Determine.
(In these steps X is the compound being determined and Y, Z the interfering ones — in this workbook the compounds are A, B.)
8. Compare your results with the sheet '_Expected results'.
Settings used in this example
• Compound: B; template: Direct (zero order, λmax)
• Measurement: Amplitude at λ — w1 = 335
• Calibration: slope 0.011976, intercept 0.000213, r = 1.00000 (n = 5)
• B absorbs alone above 300 nm, so 335 nm (its second band) is free from A.
Checks / common mistakes
Overlay the pure spectra of all compounds first: at the chosen λ every other compound must read ≈ 0. If it does not, the result is biased high.
The spectra are simulated (Gaussian bands + 0.0004 AU noise) to show the method clearly; they are not measured data. See docs/methods/README.md for the full guide with screenshots.

## Standards
Wavelength (nm) | Std: A 0, B 4 | Std: A 0, B 8 | Std: A 0, B 12 | Std: A 0, B 16 | Std: A 0, B 20
200 | 0.001334 | 0.001826 | 0.002043 | 0.003329 | 0.004122
200.5 | 0.001792 | 0.002283 | 0.00271 | 0.004424 | 0.005118
201 | 0.000256 | 0.002444 | 0.003287 | 0.004445 | 0.006072
201.5 | 0.001447 | 0.002647 | 0.004365 | 0.006269 | 0.007615
202 | 0.002165 | 0.003336 | 0.004352 | 0.007494 | 0.008696
202.5 | 0.002594 | 0.003709 | 0.006082 | 0.00803 | 0.010314
203 | 0.002647 | 0.004796 | 0.006775 | 0.009716 | 0.012266
203.5 | 0.003361 | 0.006061 | 0.008644 | 0.011149 | 0.014154
204 | 0.003302 | 0.006174 | 0.00994 | 0.012585 | 0.016118
204.5 | 0.00388 | 0.007071 | 0.011536 | 0.014293 | 0.018118
205 | 0.004259 | 0.008571 | 0.01287 | 0.015938 | 0.020636
205.5 | 0.004344 | 0.009689 | 0.014299 | 0.019657 | 0.02354
206 | 0.005081 | 0.010881 | 0.016603 | 0.021659 | 0.02696
206.5 | 0.006281 | 0.011887 | 0.018991 | 0.025527 | 0.030627
207 | 0.006673 | 0.013799 | 0.020455 | 0.027952 | 0.034138
207.5 | 0.008258 | 0.015776 | 0.023491 | 0.031107 | 0.038292
208 | 0.009179 | 0.016787 | 0.025969 | 0.034385 | 0.043433
208.5 | 0.010365 | 0.019107 | 0.029511 | 0.039137 | 0.047484
209 | 0.010687 | 0.020635 | 0.031758 | 0.042588 | 0.053417
209.5 | 0.012372 | 0.023122 | 0.035336 | 0.047904 | 0.059265
210 | 0.012645 | 0.025278 | 0.039166 | 0.051869 | 0.064641
210.5 | 0.013932 | 0.028424 | 0.04288 | 0.05711 | 0.07085
211 | 0.015603 | 0.030634 | 0.046056 | 0.061673 | 0.077548
211.5 | 0.017049 | 0.033982 | 0.050954 | 0.068684 | 0.084942
212 | 0.017711 | 0.036765 | 0.055005 | 0.073366 | 0.091703
212.5 | 0.019114 | 0.039537 | 0.059322 | 0.07907 | 0.098283
213 | 0.021436 | 0.041581 | 0.063819 | 0.085499 | 0.1063
213.5 | 0.022459 | 0.045393 | 0.0685 | 0.09086 | 0.11443
214 | 0.02481 | 0.048635 | 0.072255 | 0.097615 | 0.12203
214.5 | 0.026009 | 0.051747 | 0.078174 | 0.10294 | 0.12951
215 | 0.027487 | 0.054116 | 0.081748 | 0.10967 | 0.13673
215.5 | 0.028937 | 0.057516 | 0.085665 | 0.115 | 0.14353
216 | 0.030665 | 0.060222 | 0.090887 | 0.12049 | 0.1516
216.5 | 0.031649 | 0.063096 | 0.095713 | 0.1266 | 0.15848
217 | 0.033484 | 0.066525 | 0.098943 | 0.13162 | 0.16528
217.5 | 0.034104 | 0.068306 | 0.10356 | 0.13748 | 0.17219
218 | 0.0354 | 0.070967 | 0.10709 | 0.14197 | 0.17752
218.5 | 0.036293 | 0.073485 | 0.10935 | 0.14647 | 0.18294
219 | 0.037768 | 0.074825 | 0.1116 | 0.15015 | 0.18789
219.5 | 0.038386 | 0.076698 | 0.11462 | 0.15334 | 0.19175
220 | 0.038738 | 0.07821 | 0.11772 | 0.15634 | 0.19515
220.5 | 0.040458 | 0.07934 | 0.11865 | 0.15905 | 0.19799
221 | 0.039943 | 0.079687 | 0.11966 | 0.16028 | 0.20075
221.5 | 0.040326 | 0.080775 | 0.12058 | 0.16118 | 0.20206
222 | 0.040665 | 0.081557 | 0.12152 | 0.16164 | 0.20234
222.5 | 0.041191 | 0.080364 | 0.12146 | 0.16188 | 0.20292
223 | 0.040269 | 0.081 | 0.12133 | 0.16101 | 0.20171
223.5 | 0.040301 | 0.080087 | 0.12046 | 0.16083 | 0.19956
224 | 0.039611 | 0.078887 | 0.11834 | 0.15782 | 0.1981
224.5 | 0.038405 | 0.078782 | 0.11691 | 0.15553 | 0.1949
225 | 0.038603 | 0.076876 | 0.1144 | 0.15302 | 0.1909
225.5 | 0.037875 | 0.074453 | 0.11243 | 0.15013 | 0.18754
226 | 0.036448 | 0.073118 | 0.10917 | 0.14607 | 0.18258
226.5 | 0.035525 | 0.071293 | 0.10611 | 0.14174 | 0.1775
227 | 0.03384 | 0.068806 | 0.10417 | 0.13744 | 0.17244
227.5 | 0.032678 | 0.066847 | 0.099874 | 0.13301 | 0.16556
228 | 0.031928 | 0.064142 | 0.09651 | 0.12873 | 0.15991
228.5 | 0.031097 | 0.061961 | 0.093556 | 0.123 | 0.15407
229 | 0.029402 | 0.059756 | 0.089085 | 0.11832 | 0.14805
229.5 | 0.028445 | 0.056387 | 0.08494 | 0.11307 | 0.14138
230 | 0.027463 | 0.054233 | 0.081249 | 0.10843 | 0.13498
230.5 | 0.025884 | 0.051507 | 0.07724 | 0.10354 | 0.12881
231 | 0.024882 | 0.049355 | 0.074394 | 0.098022 | 0.12321
231.5 | 0.024085 | 0.046537 | 0.070385 | 0.094978 | 0.1178
232 | 0.021921 | 0.044606 | 0.067962 | 0.089752 | 0.11152
232.5 | 0.021699 | 0.042725 | 0.064067 | 0.085403 | 0.10618
233 | 0.019709 | 0.041286 | 0.061248 | 0.081495 | 0.1022
233.5 | 0.019784 | 0.039681 | 0.058667 | 0.078343 | 0.097906
234 | 0.018935 | 0.037179 | 0.056813 | 0.075429 | 0.093907
234.5 | 0.017979 | 0.036075 | 0.054664 | 0.073024 | 0.091227
235 | 0.017213 | 0.035547 | 0.05348 | 0.070578 | 0.088541
235.5 | 0.017175 | 0.033603 | 0.051379 | 0.068994 | 0.085774
236 | 0.016704 | 0.033551 | 0.05072 | 0.066828 | 0.084202
236.5 | 0.016689 | 0.033261 | 0.050725 | 0.066299 | 0.083477
237 | 0.016184 | 0.033597 | 0.049545 | 0.065785 | 0.082563
237.5 | 0.016778 | 0.033395 | 0.049892 | 0.066305 | 0.082238
238 | 0.016274 | 0.033246 | 0.050694 | 0.066937 | 0.083415
238.5 | 0.01681 | 0.033716 | 0.051017 | 0.067141 | 0.084298
239 | 0.017133 | 0.034479 | 0.051614 | 0.069091 | 0.086577
239.5 | 0.018267 | 0.036128 | 0.052453 | 0.071328 | 0.088595
240 | 0.017742 | 0.036509 | 0.055977 | 0.073726 | 0.091624
240.5 | 0.018863 | 0.037936 | 0.057259 | 0.07567 | 0.095353
241 | 0.020297 | 0.039789 | 0.059977 | 0.079522 | 0.099071
241.5 | 0.02046 | 0.041398 | 0.061979 | 0.083028 | 0.10336
242 | 0.021227 | 0.043367 | 0.065897 | 0.086718 | 0.10821
242.5 | 0.022935 | 0.045197 | 0.068672 | 0.092071 | 0.1148
243 | 0.023899 | 0.048026 | 0.072066 | 0.095614 | 0.12052
243.5 | 0.025636 | 0.050427 | 0.075546 | 0.1011 | 0.12591
244 | 0.026311 | 0.053825 | 0.080185 | 0.1069 | 0.13298
244.5 | 0.02822 | 0.056548 | 0.084235 | 0.11267 | 0.14095
245 | 0.029958 | 0.059373 | 0.088985 | 0.1187 | 0.14822
245.5 | 0.031735 | 0.062908 | 0.093312 | 0.12583 | 0.15661
246 | 0.032585 | 0.06592 | 0.098887 | 0.13269 | 0.16547
246.5 | 0.03538 | 0.070041 | 0.10449 | 0.13961 | 0.1741
247 | 0.036304 | 0.073326 | 0.10949 | 0.14706 | 0.18327
247.5 | 0.037981 | 0.077406 | 0.11579 | 0.15432 | 0.19274
248 | 0.040192 | 0.080631 | 0.12199 | 0.16183 | 0.20283
248.5 | 0.042684 | 0.085383 | 0.12774 | 0.17136 | 0.21326
249 | 0.044599 | 0.08878 | 0.13393 | 0.17853 | 0.22433
249.5 | 0.047063 | 0.092959 | 0.14041 | 0.18844 | 0.23486
250 | 0.049293 | 0.098068 | 0.14766 | 0.19633 | 0.24539
250.5 | 0.051475 | 0.10233 | 0.15343 | 0.20548 | 0.2571
251 | 0.053538 | 0.10735 | 0.1608 | 0.21449 | 0.26825
251.5 | 0.056694 | 0.11283 | 0.16774 | 0.22426 | 0.27925
252 | 0.057439 | 0.11632 | 0.17503 | 0.2331 | 0.29159
252.5 | 0.060014 | 0.12117 | 0.1821 | 0.24357 | 0.30394
253 | 0.062907 | 0.1263 | 0.18888 | 0.25265 | 0.31531
253.5 | 0.065022 | 0.13124 | 0.19739 | 0.26216 | 0.32799
254 | 0.066923 | 0.13582 | 0.20391 | 0.27155 | 0.3396
254.5 | 0.070049 | 0.14065 | 0.21056 | 0.28158 | 0.35185
255 | 0.072484 | 0.14585 | 0.21799 | 0.29113 | 0.36344
255.5 | 0.074208 | 0.15058 | 0.2256 | 0.30151 | 0.37677
256 | 0.077916 | 0.15473 | 0.23296 | 0.31067 | 0.38804
256.5 | 0.080452 | 0.15983 | 0.23926 | 0.3191 | 0.40006
257 | 0.082752 | 0.16453 | 0.24706 | 0.32965 | 0.41111
257.5 | 0.084341 | 0.16867 | 0.25308 | 0.33723 | 0.42214
258 | 0.086843 | 0.17367 | 0.26003 | 0.34775 | 0.4344
258.5 | 0.088521 | 0.1776 | 0.26725 | 0.35512 | 0.44456
259 | 0.091176 | 0.18257 | 0.27396 | 0.36437 | 0.45613
259.5 | 0.09288 | 0.18638 | 0.27941 | 0.37292 | 0.46568
260 | 0.095294 | 0.1903 | 0.28527 | 0.38052 | 0.4763
260.5 | 0.096893 | 0.19382 | 0.2916 | 0.38768 | 0.48486
261 | 0.099255 | 0.19764 | 0.29609 | 0.39586 | 0.4935
261.5 | 0.10089 | 0.20032 | 0.30179 | 0.40152 | 0.50307
262 | 0.10267 | 0.2039 | 0.30662 | 0.40811 | 0.51084
262.5 | 0.10304 | 0.20751 | 0.31146 | 0.41529 | 0.51842
263 | 0.10435 | 0.20931 | 0.31537 | 0.41998 | 0.52583
263.5 | 0.10652 | 0.21307 | 0.31898 | 0.42555 | 0.53146
264 | 0.10729 | 0.21435 | 0.32282 | 0.43042 | 0.5371
264.5 | 0.10872 | 0.21742 | 0.32522 | 0.43419 | 0.54232
265 | 0.10958 | 0.21858 | 0.32871 | 0.43796 | 0.54663
265.5 | 0.1103 | 0.22078 | 0.33043 | 0.44098 | 0.55044
266 | 0.11124 | 0.22179 | 0.33224 | 0.44355 | 0.55382
266.5 | 0.11159 | 0.22211 | 0.33394 | 0.44593 | 0.55643
267 | 0.11191 | 0.22394 | 0.33518 | 0.44726 | 0.55842
267.5 | 0.11158 | 0.22445 | 0.33609 | 0.44787 | 0.55919
268 | 0.1127 | 0.22399 | 0.33594 | 0.4484 | 0.55998
268.5 | 0.11199 | 0.22376 | 0.33611 | 0.44753 | 0.55998
269 | 0.11217 | 0.22339 | 0.33556 | 0.44659 | 0.55852
269.5 | 0.1115 | 0.22235 | 0.33407 | 0.44508 | 0.557
270 | 0.11138 | 0.22219 | 0.33223 | 0.44382 | 0.55399
270.5 | 0.1103 | 0.22027 | 0.3311 | 0.44162 | 0.55131
271 | 0.10956 | 0.21893 | 0.32853 | 0.43793 | 0.54753
271.5 | 0.10863 | 0.21683 | 0.32487 | 0.4335 | 0.54258
272 | 0.10726 | 0.21483 | 0.32223 | 0.42996 | 0.53706
272.5 | 0.10616 | 0.21226 | 0.31921 | 0.42573 | 0.53181
273 | 0.10505 | 0.21072 | 0.31477 | 0.42055 | 0.52548
273.5 | 0.10366 | 0.20736 | 0.31086 | 0.41489 | 0.51908
274 | 0.10181 | 0.2048 | 0.30593 | 0.40922 | 0.51081
274.5 | 0.10154 | 0.20119 | 0.30194 | 0.40275 | 0.50357
275 | 0.09849 | 0.19749 | 0.29674 | 0.39499 | 0.49339
275.5 | 0.096954 | 0.19402 | 0.2913 | 0.38848 | 0.48519
276 | 0.096161 | 0.19014 | 0.28613 | 0.38043 | 0.47565
276.5 | 0.092944 | 0.18642 | 0.27948 | 0.37293 | 0.46582
277 | 0.091465 | 0.18217 | 0.27336 | 0.36513 | 0.45551
277.5 | 0.089307 | 0.17797 | 0.26794 | 0.3569 | 0.44512
278 | 0.086783 | 0.17318 | 0.26008 | 0.34774 | 0.43491
278.5 | 0.084065 | 0.16895 | 0.2542 | 0.33834 | 0.42305
279 | 0.081739 | 0.16538 | 0.24655 | 0.3293 | 0.4122
279.5 | 0.080042 | 0.15982 | 0.24023 | 0.32039 | 0.40011
280 | 0.077867 | 0.15498 | 0.23394 | 0.3104 | 0.38846
280.5 | 0.074464 | 0.15079 | 0.22609 | 0.30128 | 0.37561
281 | 0.072998 | 0.14644 | 0.2185 | 0.29199 | 0.36464
281.5 | 0.070274 | 0.14049 | 0.21256 | 0.28251 | 0.35284
282 | 0.068488 | 0.13618 | 0.20417 | 0.2724 | 0.34012
282.5 | 0.065797 | 0.1312 | 0.1971 | 0.26305 | 0.32831
283 | 0.063011 | 0.12596 | 0.19012 | 0.25387 | 0.31624
283.5 | 0.060781 | 0.12219 | 0.18231 | 0.24392 | 0.30551
284 | 0.058608 | 0.11717 | 0.17621 | 0.23367 | 0.29248
284.5 | 0.056793 | 0.11253 | 0.16941 | 0.2251 | 0.2809
285 | 0.055168 | 0.10759 | 0.16215 | 0.21589 | 0.26859
285.5 | 0.051825 | 0.10354 | 0.15517 | 0.20654 | 0.25768
286 | 0.0495 | 0.098681 | 0.14829 | 0.19781 | 0.24724
286.5 | 0.047548 | 0.094469 | 0.14214 | 0.18846 | 0.23573
287 | 0.045145 | 0.089809 | 0.13412 | 0.181 | 0.22446
287.5 | 0.043326 | 0.085597 | 0.12907 | 0.17233 | 0.21501
288 | 0.0412 | 0.082558 | 0.12314 | 0.16383 | 0.20502
288.5 | 0.039076 | 0.077879 | 0.11633 | 0.15626 | 0.19523
289 | 0.036848 | 0.074378 | 0.11156 | 0.14858 | 0.18571
289.5 | 0.035384 | 0.070555 | 0.1058 | 0.14079 | 0.1761
290 | 0.033901 | 0.06683 | 0.1005 | 0.13295 | 0.16853
290.5 | 0.031856 | 0.063375 | 0.095408 | 0.12602 | 0.15864
291 | 0.030787 | 0.06054 | 0.090662 | 0.12138 | 0.15081
291.5 | 0.028984 | 0.057016 | 0.085889 | 0.11383 | 0.14225
292 | 0.02705 | 0.054068 | 0.080778 | 0.10814 | 0.13499
292.5 | 0.026366 | 0.051272 | 0.076847 | 0.10275 | 0.12747
293 | 0.02386 | 0.049012 | 0.072739 | 0.096938 | 0.12135
293.5 | 0.02349 | 0.046236 | 0.068962 | 0.092405 | 0.11467
294 | 0.022549 | 0.043685 | 0.065061 | 0.086684 | 0.10835
294.5 | 0.020562 | 0.041018 | 0.062254 | 0.082713 | 0.10223
295 | 0.01955 | 0.038547 | 0.059171 | 0.077364 | 0.09788
295.5 | 0.018966 | 0.037477 | 0.056255 | 0.074481 | 0.092982
296 | 0.018234 | 0.035623 | 0.053617 | 0.070749 | 0.088338
296.5 | 0.016979 | 0.03346 | 0.050871 | 0.067179 | 0.083672
297 | 0.016108 | 0.03215 | 0.047715 | 0.063986 | 0.080024
297.5 | 0.015656 | 0.030798 | 0.045483 | 0.061079 | 0.076493
298 | 0.014738 | 0.029505 | 0.0436 | 0.058691 | 0.073053
298.5 | 0.01356 | 0.028359 | 0.04221 | 0.056367 | 0.070465
299 | 0.013441 | 0.026756 | 0.040355 | 0.053498 | 0.067204
299.5 | 0.013047 | 0.026617 | 0.039414 | 0.052446 | 0.065035
300 | 0.012052 | 0.02471 | 0.038473 | 0.050038 | 0.062348
300.5 | 0.012471 | 0.024871 | 0.036618 | 0.049119 | 0.062072
301 | 0.011732 | 0.024161 | 0.036337 | 0.047825 | 0.059543
301.5 | 0.012102 | 0.023864 | 0.035115 | 0.047222 | 0.058849
302 | 0.011041 | 0.023931 | 0.034262 | 0.046226 | 0.058133
302.5 | 0.011159 | 0.023547 | 0.033922 | 0.046068 | 0.056999
303 | 0.010881 | 0.022376 | 0.034123 | 0.045477 | 0.057007
303.5 | 0.01197 | 0.022143 | 0.033469 | 0.045254 | 0.05726
304 | 0.012009 | 0.023232 | 0.034557 | 0.046237 | 0.057167
304.5 | 0.012179 | 0.022683 | 0.035238 | 0.046407 | 0.057546
305 | 0.012322 | 0.023364 | 0.034587 | 0.046943 | 0.058241
305.5 | 0.012279 | 0.024077 | 0.035478 | 0.047865 | 0.059018
306 | 0.012506 | 0.023546 | 0.036922 | 0.047983 | 0.060388
306.5 | 0.011744 | 0.023909 | 0.037268 | 0.049871 | 0.061975
307 | 0.012872 | 0.02549 | 0.038158 | 0.050351 | 0.063596
307.5 | 0.012898 | 0.02612 | 0.039049 | 0.052357 | 0.065639
308 | 0.013367 | 0.026799 | 0.040753 | 0.053722 | 0.067214
308.5 | 0.01344 | 0.027783 | 0.041783 | 0.055644 | 0.069169
309 | 0.014929 | 0.028368 | 0.042635 | 0.057703 | 0.072067
309.5 | 0.014975 | 0.029122 | 0.043618 | 0.059463 | 0.074412
310 | 0.014858 | 0.030744 | 0.045143 | 0.061253 | 0.077378
310.5 | 0.01623 | 0.03157 | 0.048202 | 0.06429 | 0.079929
311 | 0.016597 | 0.032513 | 0.049649 | 0.066811 | 0.083094
311.5 | 0.016842 | 0.034643 | 0.05224 | 0.069155 | 0.086215
312 | 0.017798 | 0.035743 | 0.053855 | 0.07195 | 0.089668
312.5 | 0.018294 | 0.037311 | 0.055214 | 0.073592 | 0.092446
313 | 0.018397 | 0.038185 | 0.057987 | 0.077189 | 0.09637
313.5 | 0.01944 | 0.039796 | 0.060566 | 0.080074 | 0.10018
314 | 0.02101 | 0.041184 | 0.062014 | 0.082385 | 0.1046
314.5 | 0.021723 | 0.042794 | 0.064519 | 0.086341 | 0.1078
315 | 0.022492 | 0.04483 | 0.067824 | 0.089505 | 0.11202
315.5 | 0.022728 | 0.046076 | 0.068886 | 0.092176 | 0.11586
316 | 0.023791 | 0.048201 | 0.071651 | 0.095781 | 0.11929
316.5 | 0.024904 | 0.049891 | 0.075052 | 0.099544 | 0.12484
317 | 0.026014 | 0.052285 | 0.077292 | 0.10319 | 0.12892
317.5 | 0.026271 | 0.052659 | 0.079367 | 0.10672 | 0.1335
318 | 0.027063 | 0.055328 | 0.081877 | 0.1095 | 0.13686
318.5 | 0.028827 | 0.056707 | 0.0853 | 0.11306 | 0.14193
319 | 0.029287 | 0.058516 | 0.088267 | 0.11684 | 0.14633
319.5 | 0.03 | 0.059724 | 0.090787 | 0.12059 | 0.1501
320 | 0.030748 | 0.061903 | 0.093296 | 0.12394 | 0.15537
320.5 | 0.031671 | 0.064165 | 0.096399 | 0.12766 | 0.15939
321 | 0.033454 | 0.065606 | 0.098584 | 0.13156 | 0.16426
321.5 | 0.034087 | 0.067432 | 0.10077 | 0.13507 | 0.1678
322 | 0.034404 | 0.069027 | 0.10453 | 0.13834 | 0.17247
322.5 | 0.034931 | 0.071328 | 0.10581 | 0.14128 | 0.17685
323 | 0.036952 | 0.072556 | 0.10879 | 0.14493 | 0.18077
323.5 | 0.037745 | 0.073353 | 0.11132 | 0.14869 | 0.18557
324 | 0.038416 | 0.075592 | 0.11438 | 0.15209 | 0.19015
324.5 | 0.03907 | 0.07668 | 0.1159 | 0.15536 | 0.19362
325 | 0.038757 | 0.077723 | 0.11843 | 0.15762 | 0.19777
325.5 | 0.039751 | 0.080322 | 0.12009 | 0.16088 | 0.20188
326 | 0.041223 | 0.082528 | 0.12313 | 0.16351 | 0.20517
326.5 | 0.041587 | 0.083421 | 0.12528 | 0.16563 | 0.20832
327 | 0.042478 | 0.084282 | 0.12739 | 0.16906 | 0.21108
327.5 | 0.042809 | 0.085662 | 0.12869 | 0.17137 | 0.2152
328 | 0.043368 | 0.087714 | 0.13099 | 0.17508 | 0.21799
328.5 | 0.044429 | 0.088479 | 0.13227 | 0.17702 | 0.22082
329 | 0.044589 | 0.089518 | 0.13406 | 0.17874 | 0.224
329.5 | 0.045252 | 0.090485 | 0.13573 | 0.18104 | 0.22645
330 | 0.045821 | 0.091453 | 0.13767 | 0.18304 | 0.22794
330.5 | 0.046237 | 0.09261 | 0.13826 | 0.18528 | 0.23078
331 | 0.046991 | 0.093276 | 0.13965 | 0.18642 | 0.23248
331.5 | 0.047416 | 0.093824 | 0.14027 | 0.18809 | 0.23465
332 | 0.046619 | 0.093917 | 0.14127 | 0.18886 | 0.23579
332.5 | 0.047313 | 0.095045 | 0.14198 | 0.1902 | 0.23744
333 | 0.047369 | 0.094984 | 0.14282 | 0.19109 | 0.23785
333.5 | 0.048093 | 0.096021 | 0.14368 | 0.19154 | 0.23877
334 | 0.04748 | 0.095308 | 0.14328 | 0.19165 | 0.23929
334.5 | 0.048287 | 0.095901 | 0.14422 | 0.19235 | 0.23978
335 | 0.048281 | 0.095999 | 0.14362 | 0.19185 | 0.23988
335.5 | 0.048208 | 0.095425 | 0.14406 | 0.19248 | 0.24033
336 | 0.047517 | 0.096503 | 0.1435 | 0.19152 | 0.23911
336.5 | 0.047703 | 0.096165 | 0.14325 | 0.19089 | 0.23895
337 | 0.047194 | 0.095069 | 0.14272 | 0.19126 | 0.23764
337.5 | 0.047576 | 0.095145 | 0.14269 | 0.18951 | 0.23649
338 | 0.047545 | 0.09448 | 0.14159 | 0.18872 | 0.23555
338.5 | 0.046256 | 0.092686 | 0.14102 | 0.18704 | 0.2337
339 | 0.046191 | 0.093147 | 0.13942 | 0.18582 | 0.23265
339.5 | 0.046072 | 0.092253 | 0.13831 | 0.18442 | 0.23037
340 | 0.045499 | 0.091459 | 0.13778 | 0.18246 | 0.22832
340.5 | 0.045842 | 0.090062 | 0.13507 | 0.18148 | 0.22638
341 | 0.044682 | 0.089374 | 0.13386 | 0.17868 | 0.22429
341.5 | 0.044298 | 0.088325 | 0.13251 | 0.17629 | 0.22075
342 | 0.043173 | 0.087714 | 0.13158 | 0.17493 | 0.21729
342.5 | 0.043622 | 0.086146 | 0.12876 | 0.17231 | 0.21484
343 | 0.042425 | 0.084718 | 0.12739 | 0.16851 | 0.21177
343.5 | 0.041991 | 0.083977 | 0.12567 | 0.16651 | 0.20883
344 | 0.04101 | 0.081731 | 0.12283 | 0.16431 | 0.20483
344.5 | 0.039393 | 0.08033 | 0.12065 | 0.16067 | 0.20083
345 | 0.038843 | 0.078241 | 0.11911 | 0.15783 | 0.19776
345.5 | 0.038185 | 0.07803 | 0.11587 | 0.15471 | 0.19391
346 | 0.038085 | 0.07618 | 0.11326 | 0.15069 | 0.18938
346.5 | 0.036869 | 0.074514 | 0.11154 | 0.14794 | 0.18546
347 | 0.036403 | 0.072732 | 0.10848 | 0.14524 | 0.18128
347.5 | 0.03549 | 0.070795 | 0.10664 | 0.14209 | 0.17788
348 | 0.034605 | 0.069098 | 0.10372 | 0.138 | 0.17227
348.5 | 0.033991 | 0.067147 | 0.10016 | 0.13446 | 0.16794
349 | 0.032805 | 0.065385 | 0.098301 | 0.13116 | 0.16361
349.5 | 0.031177 | 0.063691 | 0.095498 | 0.12741 | 0.15944
350 | 0.030658 | 0.062466 | 0.092516 | 0.12393 | 0.15452
350.5 | 0.030687 | 0.060268 | 0.0905 | 0.11975 | 0.1503
351 | 0.028684 | 0.058204 | 0.086621 | 0.11647 | 0.14591
351.5 | 0.028749 | 0.056176 | 0.0841 | 0.11303 | 0.14179
352 | 0.027957 | 0.054338 | 0.081679 | 0.10943 | 0.1364
352.5 | 0.026154 | 0.053425 | 0.079482 | 0.10495 | 0.13182
353 | 0.02563 | 0.051188 | 0.07585 | 0.10231 | 0.12732
353.5 | 0.024039 | 0.049227 | 0.073533 | 0.098454 | 0.12225
354 | 0.024024 | 0.047292 | 0.070848 | 0.094932 | 0.1181
354.5 | 0.022529 | 0.045237 | 0.067904 | 0.090661 | 0.11425
355 | 0.021452 | 0.043925 | 0.06576 | 0.087928 | 0.10993
355.5 | 0.020967 | 0.042597 | 0.062897 | 0.083871 | 0.10663
356 | 0.020102 | 0.040413 | 0.060659 | 0.080138 | 0.1019
356.5 | 0.019226 | 0.038829 | 0.05848 | 0.077887 | 0.097593
357 | 0.018729 | 0.037213 | 0.055571 | 0.07461 | 0.093823
357.5 | 0.01779 | 0.035759 | 0.053207 | 0.071463 | 0.089017
358 | 0.016816 | 0.033526 | 0.050832 | 0.068406 | 0.085169
358.5 | 0.016703 | 0.032552 | 0.048612 | 0.065303 | 0.081306
359 | 0.016552 | 0.030825 | 0.046104 | 0.062784 | 0.077839
359.5 | 0.014888 | 0.030079 | 0.044097 | 0.059786 | 0.075153
360 | 0.013879 | 0.028014 | 0.042514 | 0.057307 | 0.070688
360.5 | 0.013083 | 0.02719 | 0.040069 | 0.054943 | 0.06713
361 | 0.013054 | 0.026247 | 0.039749 | 0.051598 | 0.064096
361.5 | 0.012075 | 0.02423 | 0.03619 | 0.049047 | 0.061209
362 | 0.011564 | 0.022875 | 0.03484 | 0.047008 | 0.057968
362.5 | 0.010161 | 0.021994 | 0.032256 | 0.043503 | 0.055122
363 | 0.009928 | 0.020761 | 0.031017 | 0.04141 | 0.051415
363.5 | 0.009359 | 0.01925 | 0.029234 | 0.039414 | 0.048764
364 | 0.009665 | 0.018759 | 0.027813 | 0.036543 | 0.046334
364.5 | 0.008663 | 0.018349 | 0.026832 | 0.03549 | 0.043885
365 | 0.007282 | 0.016449 | 0.024745 | 0.033755 | 0.041914
365.5 | 0.007994 | 0.015522 | 0.022904 | 0.03072 | 0.038496
366 | 0.0078 | 0.014275 | 0.022073 | 0.029079 | 0.036981
366.5 | 0.007433 | 0.013951 | 0.020334 | 0.028024 | 0.034459
367 | 0.005375 | 0.012493 | 0.019427 | 0.026542 | 0.032717
367.5 | 0.006161 | 0.011756 | 0.018928 | 0.02375 | 0.03023
368 | 0.006046 | 0.011955 | 0.017537 | 0.023117 | 0.028048
368.5 | 0.005604 | 0.010361 | 0.016512 | 0.021376 | 0.027272
369 | 0.004925 | 0.010472 | 0.015268 | 0.01935 | 0.025208
369.5 | 0.003974 | 0.01 | 0.014072 | 0.018619 | 0.022697
370 | 0.004213 | 0.008878 | 0.012538 | 0.017327 | 0.022205
370.5 | 0.003966 | 0.008411 | 0.012091 | 0.016428 | 0.020069
371 | 0.004048 | 0.008419 | 0.01124 | 0.015295 | 0.019046
371.5 | 0.003814 | 0.007037 | 0.010944 | 0.014125 | 0.017776
372 | 0.003214 | 0.006386 | 0.010332 | 0.013195 | 0.016015
372.5 | 0.003418 | 0.005617 | 0.009429 | 0.012256 | 0.015829
373 | 0.002438 | 0.005737 | 0.008261 | 0.011648 | 0.014086
373.5 | 0.002381 | 0.0059 | 0.008703 | 0.011006 | 0.013911
374 | 0.002448 | 0.005305 | 0.00795 | 0.010229 | 0.012489
374.5 | 0.002509 | 0.004182 | 0.006751 | 0.008762 | 0.011593
375 | 0.00198 | 0.003876 | 0.006406 | 0.00812 | 0.010644
375.5 | 0.002442 | 0.003697 | 0.006105 | 0.00753 | 0.010524
376 | 0.001965 | 0.003718 | 0.005443 | 0.007771 | 0.008588
376.5 | 0.001281 | 0.00324 | 0.004944 | 0.006835 | 0.007988
377 | 0.001565 | 0.003148 | 0.004791 | 0.005817 | 0.007385
377.5 | 0.001055 | 0.002938 | 0.004233 | 0.005176 | 0.0075
378 | 0.000912 | 0.002474 | 0.003972 | 0.004784 | 0.006896
378.5 | 0.000621 | 0.002367 | 0.002931 | 0.004778 | 0.005987
379 | 0.001653 | 0.002271 | 0.002639 | 0.003935 | 0.005572
379.5 | 0.001367 | 0.001973 | 0.003707 | 0.004011 | 0.004357
380 | 0.000279 | 0.00204 | 0.00215 | 0.004217 | 0.005184
380.5 | 0.001287 | 0.002432 | 0.00272 | 0.003264 | 0.004116
381 | 0.001154 | 0.001776 | 0.002438 | 0.003421 | 0.003892
381.5 | 0.000558 | 0.001429 | 0.002199 | 0.002205 | 0.003392
382 | 0.000845 | 0.000609 | 0.001648 | 0.00271 | 0.003466
382.5 | 0.001483 | 0.001326 | 0.002162 | 0.002458 | 0.002837
383 | 0.000648 | 0.000288 | 0.00174 | 0.001764 | 0.00271
383.5 | 0.000654 | 0.000407 | 0.001319 | 0.002493 | 0.002458
384 | 0.00089 | 0.001224 | 0.000761 | 0.001762 | 0.002
384.5 | -1.3e-05 | 0.001084 | 0.00111 | 0.001327 | 0.001757
385 | -0.000143 | 0.000667 | 0.000851 | 0.001337 | 0.001972
385.5 | -0.000234 | -2.5e-05 | 0.000486 | 0.001699 | 0.001799
386 | 0.000268 | 0.000449 | 0.001048 | 0.000645 | 0.001784
386.5 | 0.000619 | 0.000269 | 0.001043 | 0.00114 | 0.001927
387 | -9.9e-05 | 0.000743 | 0.001067 | 0.000592 | 0.000726
387.5 | -0.000327 | 0.001344 | 0.000711 | 0.000972 | 0.000366
388 | 0.000286 | 0.000485 | 0.000501 | 0.000584 | 0.000831
388.5 | 3.9e-05 | 4.7e-05 | 0.000997 | 0.000687 | 0.001402
389 | 0.000413 | -0.000145 | 0.000797 | 0.00095 | 0.000787
389.5 | 0.000106 | 0.000268 | 0.001364 | 0.000597 | -2e-06
390 | 7.1e-05 | 0.00019 | 0.000622 | 0.000742 | 0.001286
390.5 | 9.9e-05 | -0.000226 | -1.6e-05 | 0.000748 | 0.000784
391 | 0.000135 | 0.000257 | 0.00071 | 0.000583 | -7.3e-05
391.5 | -0.000275 | -0.000272 | 0.001225 | 0.000467 | 0.000137
392 | -0.000348 | 0.000613 | 0.000372 | 0.000354 | 0.000459
392.5 | 0.000728 | 0.000576 | 0.000163 | 0.000556 | 0.000154
393 | 0.001321 | 0.000568 | -2.6e-05 | 0.000751 | 0.000281
393.5 | 0.000526 | -7e-05 | 0.000957 | 0.000977 | 0.000506
394 | -0.00058 | 0.000313 | 9.9e-05 | 0.000918 | 0.000268
394.5 | 7.9e-05 | 4.3e-05 | 5.9e-05 | -0.000207 | 0.000793
395 | 0.000538 | -7.1e-05 | 0.000264 | 0.000287 | 0.000135
395.5 | -0.00059 | -6e-05 | -0.000441 | 0.000347 | 0.000783
396 | -0.000141 | -0.000453 | 7e-06 | 0.00062 | -6.7e-05
396.5 | -0.000113 | -0.000518 | -0.000402 | 0.000544 | -0.000318
397 | -0.000319 | 0.00037 | 3.9e-05 | -0.000467 | -0.000706
397.5 | -0.00029 | -3e-05 | 0.000287 | 0.000204 | -0.000231
398 | 0.000448 | 0.000128 | 0.000452 | -0.000602 | 0.000139
398.5 | -0.000141 | 0.000437 | -9.4e-05 | -0.000117 | 0.000472
399 | -5.3e-05 | -0.000661 | 0.000397 | -0.000306 | -0.001292
399.5 | 0.000724 | -0.000285 | -9.5e-05 | -0.000188 | 5.3e-05
400 | -0.000305 | 9.5e-05 | -0.000605 | 0.000683 | -0.000327

## Mixtures
Wavelength (nm) | Mix: A 6, B 10 | Mix: A 10, B 6 | Mix: A 14, B 14 | Mix: A 8, B 16 | Mix: A 18, B 5
200 | 0.00959 | 0.014284 | 0.020531 | 0.012826 | 0.022835
200.5 | 0.011623 | 0.014559 | 0.023229 | 0.015286 | 0.027107
201 | 0.013324 | 0.01869 | 0.026788 | 0.018378 | 0.031516
201.5 | 0.015364 | 0.021148 | 0.0305 | 0.020891 | 0.036251
202 | 0.01801 | 0.024774 | 0.036267 | 0.024533 | 0.04182
202.5 | 0.019423 | 0.027855 | 0.042154 | 0.028231 | 0.046731
203 | 0.022463 | 0.032225 | 0.047445 | 0.031298 | 0.053447
203.5 | 0.025271 | 0.035646 | 0.053958 | 0.036648 | 0.059988
204 | 0.029127 | 0.039916 | 0.06104 | 0.040779 | 0.067031
204.5 | 0.033231 | 0.045192 | 0.067715 | 0.047052 | 0.07534
205 | 0.036776 | 0.050313 | 0.07712 | 0.052129 | 0.084235
205.5 | 0.041226 | 0.055509 | 0.084244 | 0.057671 | 0.092885
206 | 0.044793 | 0.061377 | 0.093803 | 0.064436 | 0.10272
206.5 | 0.05042 | 0.067878 | 0.1043 | 0.071321 | 0.11322
207 | 0.055065 | 0.074247 | 0.11317 | 0.078639 | 0.12335
207.5 | 0.061448 | 0.080367 | 0.12369 | 0.086707 | 0.13401
208 | 0.066623 | 0.087202 | 0.13421 | 0.094346 | 0.14481
208.5 | 0.072374 | 0.093548 | 0.14562 | 0.10309 | 0.15563
209 | 0.077861 | 0.10195 | 0.15592 | 0.11123 | 0.16604
209.5 | 0.084149 | 0.10823 | 0.16794 | 0.11968 | 0.17824
210 | 0.090279 | 0.11588 | 0.18036 | 0.12808 | 0.18815
210.5 | 0.096042 | 0.12244 | 0.19123 | 0.13712 | 0.19951
211 | 0.10216 | 0.12985 | 0.2024 | 0.14694 | 0.20971
211.5 | 0.10845 | 0.13628 | 0.21379 | 0.15623 | 0.21969
212 | 0.11432 | 0.14212 | 0.22465 | 0.16556 | 0.22909
212.5 | 0.12026 | 0.14861 | 0.23502 | 0.17363 | 0.23758
213 | 0.12616 | 0.1533 | 0.24519 | 0.1826 | 0.24611
213.5 | 0.13156 | 0.15944 | 0.25378 | 0.19068 | 0.25251
214 | 0.13748 | 0.16316 | 0.26355 | 0.19895 | 0.2598
214.5 | 0.14159 | 0.16824 | 0.27127 | 0.20595 | 0.26443
215 | 0.14729 | 0.17239 | 0.27937 | 0.21449 | 0.27033
215.5 | 0.15111 | 0.17541 | 0.28583 | 0.22076 | 0.27353
216 | 0.15518 | 0.17802 | 0.29147 | 0.22737 | 0.27691
216.5 | 0.15851 | 0.18021 | 0.29711 | 0.23339 | 0.27958
217 | 0.16227 | 0.18227 | 0.30123 | 0.2386 | 0.27954
217.5 | 0.1651 | 0.18425 | 0.30581 | 0.24262 | 0.28101
218 | 0.16776 | 0.18519 | 0.30844 | 0.24785 | 0.28147
218.5 | 0.16957 | 0.18536 | 0.31118 | 0.25106 | 0.28117
219 | 0.17154 | 0.18527 | 0.31216 | 0.25389 | 0.2805
219.5 | 0.1736 | 0.18606 | 0.31448 | 0.25581 | 0.27903
220 | 0.17406 | 0.18503 | 0.31445 | 0.25728 | 0.27764
220.5 | 0.17439 | 0.18507 | 0.31547 | 0.25862 | 0.27606
221 | 0.17545 | 0.18459 | 0.3153 | 0.26007 | 0.27476
221.5 | 0.17522 | 0.18476 | 0.31433 | 0.26068 | 0.27353
222 | 0.17552 | 0.18421 | 0.31459 | 0.26043 | 0.27306
222.5 | 0.17536 | 0.18407 | 0.31391 | 0.26121 | 0.27234
223 | 0.17468 | 0.18396 | 0.31401 | 0.26012 | 0.27272
223.5 | 0.1737 | 0.1838 | 0.3143 | 0.25971 | 0.27394
224 | 0.17394 | 0.18482 | 0.31413 | 0.25885 | 0.27538
224.5 | 0.17392 | 0.18609 | 0.31474 | 0.25716 | 0.27726
225 | 0.17322 | 0.18684 | 0.31473 | 0.25708 | 0.28078
225.5 | 0.17332 | 0.18891 | 0.31666 | 0.25601 | 0.28555
226 | 0.17344 | 0.19092 | 0.31866 | 0.25459 | 0.2905
226.5 | 0.17292 | 0.19423 | 0.32153 | 0.25483 | 0.29662
227 | 0.17394 | 0.1975 | 0.3241 | 0.25465 | 0.30468
227.5 | 0.17465 | 0.20159 | 0.32856 | 0.25459 | 0.31418
228 | 0.17527 | 0.20598 | 0.33327 | 0.25493 | 0.32391
228.5 | 0.17645 | 0.21171 | 0.33805 | 0.25465 | 0.33502
229 | 0.1779 | 0.21672 | 0.34462 | 0.25625 | 0.34775
229.5 | 0.1802 | 0.22315 | 0.35243 | 0.25773 | 0.36041
230 | 0.18157 | 0.23005 | 0.35984 | 0.26006 | 0.37616
230.5 | 0.18407 | 0.23811 | 0.3687 | 0.2627 | 0.39069
231 | 0.18727 | 0.24574 | 0.37925 | 0.26626 | 0.4069
231.5 | 0.19002 | 0.25437 | 0.3895 | 0.26982 | 0.42506
232 | 0.1935 | 0.26404 | 0.40054 | 0.27344 | 0.44181
232.5 | 0.19788 | 0.27378 | 0.41265 | 0.27894 | 0.46074
233 | 0.20286 | 0.28352 | 0.42509 | 0.28458 | 0.48021
233.5 | 0.20773 | 0.29254 | 0.4383 | 0.28992 | 0.49968
234 | 0.21209 | 0.30385 | 0.45214 | 0.29558 | 0.52015
234.5 | 0.21743 | 0.31468 | 0.46678 | 0.30278 | 0.54104
235 | 0.22381 | 0.32569 | 0.48187 | 0.31027 | 0.5616
235.5 | 0.22987 | 0.3372 | 0.49621 | 0.3174 | 0.5823
236 | 0.23615 | 0.34936 | 0.51167 | 0.32608 | 0.60347
236.5 | 0.24285 | 0.35978 | 0.52708 | 0.33403 | 0.62443
237 | 0.24963 | 0.37081 | 0.54221 | 0.34353 | 0.64463
237.5 | 0.2559 | 0.38247 | 0.55836 | 0.3524 | 0.6631
238 | 0.2628 | 0.39312 | 0.57404 | 0.36117 | 0.68392
238.5 | 0.2689 | 0.40379 | 0.5886 | 0.3699 | 0.70249
239 | 0.27563 | 0.4138 | 0.60424 | 0.37978 | 0.72042
239.5 | 0.28292 | 0.42402 | 0.61827 | 0.38869 | 0.7371
240 | 0.28974 | 0.43372 | 0.6331 | 0.39791 | 0.75361
240.5 | 0.29628 | 0.44181 | 0.64674 | 0.40779 | 0.76897
241 | 0.3026 | 0.45042 | 0.65909 | 0.41606 | 0.78262
241.5 | 0.30885 | 0.4594 | 0.67076 | 0.42465 | 0.79658
242 | 0.31482 | 0.46656 | 0.68299 | 0.43412 | 0.80807
242.5 | 0.3195 | 0.47261 | 0.69333 | 0.44286 | 0.81797
243 | 0.32534 | 0.47828 | 0.70326 | 0.45084 | 0.82643
243.5 | 0.33096 | 0.48374 | 0.71267 | 0.45821 | 0.83472
244 | 0.33584 | 0.48771 | 0.72135 | 0.46609 | 0.84031
244.5 | 0.33948 | 0.49163 | 0.72784 | 0.47207 | 0.8448
245 | 0.3447 | 0.49521 | 0.73385 | 0.47964 | 0.84697
245.5 | 0.34785 | 0.49658 | 0.73824 | 0.48576 | 0.84929
246 | 0.35137 | 0.49816 | 0.74277 | 0.4905 | 0.8479
246.5 | 0.35464 | 0.49797 | 0.74545 | 0.49608 | 0.84681
247 | 0.35754 | 0.49777 | 0.74821 | 0.50024 | 0.84176
247.5 | 0.35901 | 0.49696 | 0.74878 | 0.50476 | 0.83771
248 | 0.36197 | 0.4939 | 0.74899 | 0.50938 | 0.83118
248.5 | 0.36326 | 0.49158 | 0.74808 | 0.51225 | 0.82229
249 | 0.36424 | 0.48824 | 0.74617 | 0.51597 | 0.81383
249.5 | 0.36543 | 0.48432 | 0.74391 | 0.51883 | 0.8031
250 | 0.36536 | 0.4801 | 0.74019 | 0.5211 | 0.79131
250.5 | 0.36638 | 0.47442 | 0.73556 | 0.52348 | 0.77871
251 | 0.36685 | 0.46766 | 0.73072 | 0.52558 | 0.76462
251.5 | 0.36638 | 0.46176 | 0.72519 | 0.52628 | 0.75033
252 | 0.36619 | 0.45465 | 0.71921 | 0.52773 | 0.73462
252.5 | 0.36571 | 0.44736 | 0.71187 | 0.52877 | 0.718
253 | 0.36505 | 0.4409 | 0.70506 | 0.5286 | 0.70033
253.5 | 0.36453 | 0.43171 | 0.69649 | 0.52892 | 0.68264
254 | 0.36305 | 0.42401 | 0.68851 | 0.52841 | 0.66519
254.5 | 0.36177 | 0.41518 | 0.67958 | 0.52958 | 0.64568
255 | 0.36078 | 0.4072 | 0.67183 | 0.52971 | 0.6268
255.5 | 0.36 | 0.39826 | 0.66221 | 0.5292 | 0.60735
256 | 0.35756 | 0.38935 | 0.65321 | 0.52949 | 0.58847
256.5 | 0.35636 | 0.38041 | 0.64486 | 0.52878 | 0.56924
257 | 0.35393 | 0.37204 | 0.63516 | 0.52792 | 0.54989
257.5 | 0.35346 | 0.36239 | 0.62592 | 0.52693 | 0.53055
258 | 0.35152 | 0.35372 | 0.61749 | 0.52586 | 0.51166
258.5 | 0.3501 | 0.34491 | 0.60799 | 0.52546 | 0.49256
259 | 0.34783 | 0.33738 | 0.59948 | 0.52475 | 0.47448
259.5 | 0.34571 | 0.32836 | 0.59012 | 0.52359 | 0.4558
260 | 0.34474 | 0.31987 | 0.58188 | 0.52249 | 0.43862
260.5 | 0.34292 | 0.31212 | 0.57348 | 0.52138 | 0.42178
261 | 0.3405 | 0.3055 | 0.56471 | 0.52065 | 0.40438
261.5 | 0.33846 | 0.29689 | 0.5572 | 0.51843 | 0.38898
262 | 0.33741 | 0.28992 | 0.5483 | 0.51764 | 0.37323
262.5 | 0.33531 | 0.28266 | 0.54053 | 0.51633 | 0.35762
263 | 0.33453 | 0.27559 | 0.53297 | 0.51499 | 0.34301
263.5 | 0.3313 | 0.26867 | 0.52583 | 0.51327 | 0.32921
264 | 0.32934 | 0.26239 | 0.51796 | 0.51101 | 0.31622
264.5 | 0.32745 | 0.25619 | 0.51068 | 0.50881 | 0.30412
265 | 0.32469 | 0.24948 | 0.50386 | 0.50633 | 0.29204
265.5 | 0.32278 | 0.24564 | 0.49727 | 0.50481 | 0.28094
266 | 0.32147 | 0.23941 | 0.48971 | 0.50191 | 0.27043
266.5 | 0.31833 | 0.23372 | 0.48254 | 0.49885 | 0.25894
267 | 0.3156 | 0.2284 | 0.47623 | 0.49532 | 0.24846
267.5 | 0.31266 | 0.22314 | 0.46979 | 0.49174 | 0.23931
268 | 0.31059 | 0.21896 | 0.4626 | 0.48896 | 0.23022
268.5 | 0.30767 | 0.21377 | 0.45614 | 0.4845 | 0.22233
269 | 0.3044 | 0.20942 | 0.44859 | 0.48041 | 0.21439
269.5 | 0.30087 | 0.20417 | 0.44254 | 0.47635 | 0.20753
270 | 0.29772 | 0.20047 | 0.43572 | 0.47011 | 0.1992
270.5 | 0.29414 | 0.1968 | 0.42899 | 0.46518 | 0.19285
271 | 0.28994 | 0.19164 | 0.42185 | 0.46009 | 0.18632
271.5 | 0.28651 | 0.18789 | 0.41491 | 0.45432 | 0.18015
272 | 0.28128 | 0.18345 | 0.40698 | 0.44746 | 0.17403
272.5 | 0.27731 | 0.17979 | 0.40029 | 0.44134 | 0.16941
273 | 0.27322 | 0.17572 | 0.39189 | 0.4343 | 0.16304
273.5 | 0.26865 | 0.17128 | 0.38471 | 0.42718 | 0.15789
274 | 0.26341 | 0.16701 | 0.37766 | 0.42013 | 0.15311
274.5 | 0.25902 | 0.16335 | 0.3697 | 0.41234 | 0.14805
275 | 0.25418 | 0.15926 | 0.36063 | 0.40475 | 0.14396
275.5 | 0.24865 | 0.155 | 0.35349 | 0.39627 | 0.1382
276 | 0.24314 | 0.15144 | 0.34484 | 0.38813 | 0.13417
276.5 | 0.23783 | 0.14744 | 0.33636 | 0.37848 | 0.12973
277 | 0.23198 | 0.14292 | 0.32829 | 0.37004 | 0.12569
277.5 | 0.22558 | 0.13924 | 0.31971 | 0.36019 | 0.12196
278 | 0.22008 | 0.13647 | 0.31159 | 0.3508 | 0.11734
278.5 | 0.21472 | 0.13164 | 0.30175 | 0.34148 | 0.11285
279 | 0.20772 | 0.12737 | 0.2934 | 0.33276 | 0.1098
279.5 | 0.20211 | 0.12436 | 0.28471 | 0.3229 | 0.10556
280 | 0.1958 | 0.11922 | 0.2752 | 0.31242 | 0.1025
280.5 | 0.18943 | 0.11446 | 0.26756 | 0.3038 | 0.098681
281 | 0.18321 | 0.11077 | 0.25896 | 0.29384 | 0.095833
281.5 | 0.17748 | 0.10702 | 0.24919 | 0.28353 | 0.091677
282 | 0.17139 | 0.10387 | 0.24081 | 0.27368 | 0.087659
282.5 | 0.16525 | 0.10033 | 0.23182 | 0.26398 | 0.084729
283 | 0.15844 | 0.095282 | 0.22338 | 0.25376 | 0.081665
283.5 | 0.15285 | 0.092857 | 0.21443 | 0.24553 | 0.077992
284 | 0.14711 | 0.088231 | 0.20657 | 0.23484 | 0.074732
284.5 | 0.14144 | 0.085195 | 0.19742 | 0.2255 | 0.071492
285 | 0.1355 | 0.081712 | 0.19021 | 0.21614 | 0.068359
285.5 | 0.12984 | 0.078202 | 0.18173 | 0.20685 | 0.065151
286 | 0.12432 | 0.074609 | 0.17465 | 0.19778 | 0.063073
286.5 | 0.11871 | 0.071328 | 0.16527 | 0.18948 | 0.059895
287 | 0.11267 | 0.067541 | 0.15833 | 0.18025 | 0.056388
287.5 | 0.10778 | 0.064932 | 0.15049 | 0.17195 | 0.054696
288 | 0.10237 | 0.062126 | 0.14379 | 0.1639 | 0.051663
288.5 | 0.096758 | 0.059113 | 0.13707 | 0.15663 | 0.049217
289 | 0.092245 | 0.056892 | 0.1298 | 0.14824 | 0.046747
289.5 | 0.089122 | 0.053485 | 0.1231 | 0.14083 | 0.044807
290 | 0.083373 | 0.05065 | 0.11671 | 0.13422 | 0.041176
290.5 | 0.078628 | 0.048436 | 0.11147 | 0.12677 | 0.039885
291 | 0.075169 | 0.045384 | 0.10612 | 0.1205 | 0.038031
291.5 | 0.072065 | 0.042396 | 0.10054 | 0.11467 | 0.035766
292 | 0.067672 | 0.039749 | 0.094807 | 0.10846 | 0.03326
292.5 | 0.064266 | 0.038998 | 0.089381 | 0.10257 | 0.031605
293 | 0.060588 | 0.036666 | 0.084848 | 0.097716 | 0.030485
293.5 | 0.057438 | 0.034839 | 0.079872 | 0.091778 | 0.029347
294 | 0.054935 | 0.032538 | 0.075581 | 0.086833 | 0.026474
294.5 | 0.051884 | 0.031089 | 0.072628 | 0.081793 | 0.025577
295 | 0.048848 | 0.030092 | 0.068333 | 0.078219 | 0.024594
295.5 | 0.046736 | 0.027969 | 0.065063 | 0.074452 | 0.023308
296 | 0.044482 | 0.025991 | 0.061568 | 0.070208 | 0.022265
296.5 | 0.041812 | 0.024972 | 0.059194 | 0.067787 | 0.021434
297 | 0.040028 | 0.02451 | 0.056466 | 0.064275 | 0.020288
297.5 | 0.038212 | 0.022434 | 0.053047 | 0.061341 | 0.018484
298 | 0.036201 | 0.020851 | 0.050911 | 0.05786 | 0.018131
298.5 | 0.035451 | 0.020079 | 0.048718 | 0.056096 | 0.017073
299 | 0.033607 | 0.01967 | 0.047672 | 0.05402 | 0.016849
299.5 | 0.032268 | 0.01983 | 0.046022 | 0.0512 | 0.016878
300 | 0.030841 | 0.018944 | 0.044001 | 0.050403 | 0.016183
300.5 | 0.030153 | 0.018094 | 0.043844 | 0.048952 | 0.016274
301 | 0.030565 | 0.017459 | 0.041948 | 0.047225 | 0.015547
301.5 | 0.028938 | 0.017591 | 0.040666 | 0.047222 | 0.015153
302 | 0.029146 | 0.017321 | 0.039751 | 0.047176 | 0.014917
302.5 | 0.028545 | 0.01699 | 0.040111 | 0.046202 | 0.014096
303 | 0.028747 | 0.017028 | 0.040089 | 0.045321 | 0.014237
303.5 | 0.028765 | 0.016809 | 0.039699 | 0.045529 | 0.014522
304 | 0.028321 | 0.0179 | 0.040229 | 0.046255 | 0.014643
304.5 | 0.028745 | 0.01637 | 0.040884 | 0.045839 | 0.014696
305 | 0.029251 | 0.017576 | 0.040871 | 0.046875 | 0.014254
305.5 | 0.029722 | 0.0174 | 0.041884 | 0.047491 | 0.014436
306 | 0.030243 | 0.018435 | 0.043022 | 0.048475 | 0.014697
306.5 | 0.030816 | 0.018471 | 0.043517 | 0.050166 | 0.016233
307 | 0.031985 | 0.019187 | 0.044195 | 0.050159 | 0.015846
307.5 | 0.033155 | 0.019789 | 0.045852 | 0.052671 | 0.01667
308 | 0.033837 | 0.020255 | 0.047284 | 0.053898 | 0.016811
308.5 | 0.034414 | 0.021222 | 0.049449 | 0.055496 | 0.017311
309 | 0.036049 | 0.020978 | 0.050073 | 0.056753 | 0.017874
309.5 | 0.037184 | 0.022676 | 0.051403 | 0.059576 | 0.018449
310 | 0.038698 | 0.02302 | 0.054038 | 0.061454 | 0.019307
310.5 | 0.040687 | 0.024122 | 0.056059 | 0.062735 | 0.020318
311 | 0.041377 | 0.0255 | 0.057881 | 0.065786 | 0.02024
311.5 | 0.043046 | 0.02628 | 0.060497 | 0.068396 | 0.021563
312 | 0.044355 | 0.027619 | 0.062721 | 0.071125 | 0.022258
312.5 | 0.04632 | 0.026961 | 0.065494 | 0.074047 | 0.023064
313 | 0.047998 | 0.028553 | 0.067181 | 0.076822 | 0.024395
313.5 | 0.050714 | 0.03043 | 0.069823 | 0.080741 | 0.025695
314 | 0.052101 | 0.030735 | 0.072762 | 0.083592 | 0.025471
314.5 | 0.053707 | 0.032288 | 0.075307 | 0.086261 | 0.027048
315 | 0.056711 | 0.034247 | 0.078824 | 0.090058 | 0.02842
315.5 | 0.057455 | 0.035142 | 0.081413 | 0.093634 | 0.029351
316 | 0.060417 | 0.035659 | 0.084093 | 0.096311 | 0.029409
316.5 | 0.0626 | 0.037498 | 0.087682 | 0.098977 | 0.031314
317 | 0.064574 | 0.038865 | 0.090282 | 0.1031 | 0.03203
317.5 | 0.066415 | 0.039939 | 0.093528 | 0.107 | 0.033102
318 | 0.068665 | 0.04119 | 0.096525 | 0.10954 | 0.034677
318.5 | 0.070791 | 0.042614 | 0.099241 | 0.11348 | 0.035073
319 | 0.073466 | 0.043291 | 0.10236 | 0.11742 | 0.036687
319.5 | 0.074449 | 0.045279 | 0.10554 | 0.12108 | 0.038048
320 | 0.077405 | 0.046179 | 0.10822 | 0.12452 | 0.038565
320.5 | 0.079895 | 0.048058 | 0.11163 | 0.12788 | 0.039537
321 | 0.081817 | 0.049401 | 0.11521 | 0.13122 | 0.041898
321.5 | 0.083522 | 0.050192 | 0.11807 | 0.13457 | 0.041741
322 | 0.086332 | 0.052102 | 0.12062 | 0.13869 | 0.043757
322.5 | 0.088281 | 0.053292 | 0.12496 | 0.1421 | 0.044036
323 | 0.091026 | 0.054406 | 0.12658 | 0.14493 | 0.045444
323.5 | 0.093351 | 0.055965 | 0.13016 | 0.14819 | 0.046213
324 | 0.094864 | 0.056777 | 0.13306 | 0.15224 | 0.047493
324.5 | 0.096061 | 0.057987 | 0.13608 | 0.15534 | 0.048497
325 | 0.098701 | 0.059299 | 0.13795 | 0.15774 | 0.04897
325.5 | 0.1003 | 0.060377 | 0.14094 | 0.1604 | 0.050288
326 | 0.10277 | 0.061953 | 0.14341 | 0.16404 | 0.051633
326.5 | 0.10493 | 0.061602 | 0.14641 | 0.16713 | 0.051829
327 | 0.10607 | 0.063439 | 0.14847 | 0.16989 | 0.052092
327.5 | 0.10836 | 0.063723 | 0.15074 | 0.1724 | 0.053895
328 | 0.10968 | 0.065236 | 0.15293 | 0.1739 | 0.054214
328.5 | 0.1108 | 0.066405 | 0.15552 | 0.17685 | 0.055416
329 | 0.11153 | 0.067115 | 0.1569 | 0.17884 | 0.055789
329.5 | 0.11322 | 0.068256 | 0.15875 | 0.18072 | 0.056262
330 | 0.11376 | 0.068693 | 0.15965 | 0.18393 | 0.057051
330.5 | 0.11582 | 0.068105 | 0.16179 | 0.18457 | 0.057829
331 | 0.11642 | 0.069328 | 0.16236 | 0.18621 | 0.058295
331.5 | 0.11777 | 0.070909 | 0.16443 | 0.18735 | 0.05807
332 | 0.11808 | 0.071357 | 0.16516 | 0.18894 | 0.058413
332.5 | 0.11783 | 0.071152 | 0.16564 | 0.19015 | 0.05878
333 | 0.11939 | 0.070802 | 0.16626 | 0.1904 | 0.059997
333.5 | 0.11991 | 0.072535 | 0.16725 | 0.19115 | 0.059809
334 | 0.12049 | 0.071001 | 0.16775 | 0.19145 | 0.059518
334.5 | 0.11961 | 0.072067 | 0.16772 | 0.1915 | 0.059861
335 | 0.11969 | 0.071957 | 0.1679 | 0.19193 | 0.060526
335.5 | 0.12003 | 0.072038 | 0.16782 | 0.19205 | 0.059512
336 | 0.1194 | 0.071364 | 0.16749 | 0.19163 | 0.059709
336.5 | 0.11898 | 0.071496 | 0.16722 | 0.19106 | 0.060032
337 | 0.11869 | 0.071841 | 0.16683 | 0.18975 | 0.059817
337.5 | 0.11964 | 0.070901 | 0.16651 | 0.1898 | 0.059523
338 | 0.11771 | 0.070161 | 0.16536 | 0.18872 | 0.058585
338.5 | 0.11681 | 0.07026 | 0.16383 | 0.18807 | 0.058574
339 | 0.11618 | 0.069788 | 0.16351 | 0.18583 | 0.057186
339.5 | 0.11539 | 0.06912 | 0.16199 | 0.18399 | 0.057513
340 | 0.11484 | 0.068218 | 0.15993 | 0.18207 | 0.056884
340.5 | 0.11354 | 0.068052 | 0.15881 | 0.18072 | 0.056315
341 | 0.11221 | 0.067026 | 0.15642 | 0.17893 | 0.056385
341.5 | 0.11055 | 0.065791 | 0.15443 | 0.17688 | 0.055156
342 | 0.10915 | 0.065482 | 0.15242 | 0.17383 | 0.054428
342.5 | 0.1073 | 0.064462 | 0.15121 | 0.17179 | 0.054023
343 | 0.10534 | 0.063563 | 0.14836 | 0.16984 | 0.053655
343.5 | 0.10506 | 0.061807 | 0.14606 | 0.16675 | 0.051758
344 | 0.1023 | 0.061122 | 0.14349 | 0.16376 | 0.051498
344.5 | 0.10105 | 0.060513 | 0.14073 | 0.16081 | 0.050172
345 | 0.098622 | 0.059566 | 0.13843 | 0.15729 | 0.04872
345.5 | 0.096756 | 0.058159 | 0.13577 | 0.15423 | 0.048138
346 | 0.09472 | 0.05688 | 0.13347 | 0.15137 | 0.04749
346.5 | 0.092171 | 0.055873 | 0.12947 | 0.14877 | 0.046232
347 | 0.090588 | 0.054716 | 0.12748 | 0.14538 | 0.045102
347.5 | 0.088629 | 0.053747 | 0.12413 | 0.14143 | 0.044205
348 | 0.086633 | 0.051925 | 0.12013 | 0.13826 | 0.043009
348.5 | 0.084295 | 0.050189 | 0.11813 | 0.13465 | 0.041907
349 | 0.082278 | 0.049273 | 0.11527 | 0.13139 | 0.040495
349.5 | 0.079741 | 0.046954 | 0.11191 | 0.12705 | 0.040017
350 | 0.07736 | 0.046404 | 0.10841 | 0.12382 | 0.037741
350.5 | 0.074306 | 0.044844 | 0.1052 | 0.12046 | 0.037593
351 | 0.072604 | 0.04365 | 0.10116 | 0.11594 | 0.036075
351.5 | 0.071166 | 0.042493 | 0.098899 | 0.11278 | 0.035555
352 | 0.068324 | 0.040956 | 0.095665 | 0.10928 | 0.033652
352.5 | 0.066333 | 0.038999 | 0.092056 | 0.10573 | 0.032041
353 | 0.063681 | 0.038088 | 0.089359 | 0.10217 | 0.032053
353.5 | 0.061264 | 0.036572 | 0.085891 | 0.09863 | 0.031148
354 | 0.060141 | 0.035559 | 0.08396 | 0.095257 | 0.028902
354.5 | 0.057278 | 0.0344 | 0.079507 | 0.091913 | 0.028706
355 | 0.055572 | 0.032573 | 0.076608 | 0.08805 | 0.027573
355.5 | 0.053055 | 0.031975 | 0.073119 | 0.085371 | 0.026004
356 | 0.050628 | 0.030945 | 0.070883 | 0.080946 | 0.025597
356.5 | 0.047786 | 0.029079 | 0.068199 | 0.077761 | 0.023974
357 | 0.046697 | 0.028265 | 0.065322 | 0.074732 | 0.023707
357.5 | 0.04405 | 0.026191 | 0.062284 | 0.071117 | 0.022793
358 | 0.041848 | 0.025403 | 0.060172 | 0.067735 | 0.021179
358.5 | 0.041337 | 0.024868 | 0.057261 | 0.064765 | 0.020263
359 | 0.03928 | 0.022666 | 0.05505 | 0.062798 | 0.019504
359.5 | 0.037446 | 0.022089 | 0.052007 | 0.059187 | 0.017809
360 | 0.035455 | 0.020881 | 0.048893 | 0.056107 | 0.017432
360.5 | 0.033717 | 0.020677 | 0.047151 | 0.054673 | 0.016843
361 | 0.032248 | 0.019684 | 0.04449 | 0.051423 | 0.01691
361.5 | 0.030539 | 0.018015 | 0.042981 | 0.048788 | 0.015697
362 | 0.028955 | 0.017384 | 0.040237 | 0.045846 | 0.014545
362.5 | 0.027725 | 0.015391 | 0.038283 | 0.043943 | 0.013345
363 | 0.025 | 0.015909 | 0.036348 | 0.041316 | 0.013121
363.5 | 0.024947 | 0.014873 | 0.033796 | 0.039389 | 0.012308
364 | 0.022995 | 0.014459 | 0.03237 | 0.037134 | 0.011064
364.5 | 0.022301 | 0.0122 | 0.031067 | 0.034989 | 0.010854
365 | 0.02119 | 0.012551 | 0.028756 | 0.033185 | 0.009799
365.5 | 0.019194 | 0.011559 | 0.027255 | 0.030838 | 0.009543
366 | 0.018711 | 0.009973 | 0.025794 | 0.029211 | 0.007868
366.5 | 0.017617 | 0.010088 | 0.023936 | 0.027323 | 0.009203
367 | 0.015977 | 0.00971 | 0.022632 | 0.026133 | 0.007903
367.5 | 0.015053 | 0.009468 | 0.021536 | 0.024674 | 0.007183
368 | 0.014363 | 0.008658 | 0.019886 | 0.023201 | 0.006882
368.5 | 0.012524 | 0.007834 | 0.018798 | 0.020668 | 0.006027
369 | 0.012379 | 0.007705 | 0.017522 | 0.020763 | 0.006191
369.5 | 0.011819 | 0.006997 | 0.016024 | 0.018685 | 0.006394
370 | 0.010794 | 0.006318 | 0.014845 | 0.017491 | 0.005585
370.5 | 0.009603 | 0.005732 | 0.015131 | 0.016353 | 0.00496
371 | 0.009395 | 0.005785 | 0.013924 | 0.015243 | 0.004702
371.5 | 0.009506 | 0.005397 | 0.012411 | 0.014725 | 0.004473
372 | 0.008813 | 0.005701 | 0.011976 | 0.01329 | 0.003969
372.5 | 0.007653 | 0.004436 | 0.011225 | 0.012428 | 0.003597
373 | 0.006831 | 0.00493 | 0.010401 | 0.011813 | 0.003568
373.5 | 0.006632 | 0.00412 | 0.009151 | 0.011136 | 0.003408
374 | 0.006161 | 0.004161 | 0.008047 | 0.009431 | 0.002842
374.5 | 0.005429 | 0.003055 | 0.00748 | 0.008623 | 0.00295
375 | 0.00514 | 0.002812 | 0.007293 | 0.007661 | 0.00214
375.5 | 0.005126 | 0.002799 | 0.007369 | 0.007031 | 0.001701
376 | 0.004877 | 0.003058 | 0.006931 | 0.007318 | 0.001471
376.5 | 0.003178 | 0.002561 | 0.005725 | 0.006676 | 0.0022
377 | 0.00422 | 0.002597 | 0.005233 | 0.006107 | 0.00159
377.5 | 0.00358 | 0.002241 | 0.005422 | 0.005769 | 0.001719
378 | 0.003704 | 0.002356 | 0.005369 | 0.005305 | 0.00129
378.5 | 0.002973 | 0.001866 | 0.003811 | 0.005242 | 0.001319
379 | 0.002913 | 0.001431 | 0.004026 | 0.005056 | 0.001415
379.5 | 0.003167 | 0.001711 | 0.00358 | 0.003692 | 0.000667
380 | 0.001112 | 0.001415 | 0.0031 | 0.004276 | 0.000868
380.5 | 0.002743 | 0.001962 | 0.002829 | 0.003794 | 0.001615
381 | 0.002002 | 0.001091 | 0.002572 | 0.002762 | 0.001718
381.5 | 0.00103 | 0.001099 | 0.00247 | 0.002889 | 0.000923
382 | 0.000539 | 0.00084 | 0.001643 | 0.0024 | 0.000862
382.5 | 0.00159 | 0.00134 | 0.002059 | 0.001669 | 0.000475
383 | 0.000923 | 0.00115 | 0.002191 | 0.002154 | 0.000221
383.5 | 0.001616 | 0.000988 | 0.002011 | 0.001297 | 0.000902
384 | 0.00119 | 0.000163 | 0.001871 | 0.001788 | 0.001021
384.5 | 0.000949 | -6e-06 | 0.001546 | 0.001192 | 0.000672
385 | 0.001308 | 0.000651 | 0.001362 | 0.001328 | 0.000658
385.5 | 0.000903 | 0.000783 | 0.001942 | 0.001646 | 0.000112
386 | 0.001044 | 0.001039 | 0.000986 | 0.001247 | 0.000828
386.5 | 0.001009 | -0.000313 | 0.001022 | 0.000766 | 0.00044
387 | 0.00074 | 0.000298 | 0.001365 | 0.001073 | 5.2e-05
387.5 | 0.000868 | -0.000346 | 0.0008 | 0.001163 | 0.000694
388 | 0.000814 | -8.6e-05 | 0.000193 | 0.001187 | -8.7e-05
388.5 | -9.9e-05 | 0.00018 | 0.00042 | 0.000391 | -0.000173
389 | 0.000295 | 0.000627 | 0.000406 | 0.001102 | 0.000248
389.5 | 0.000315 | 7e-06 | 0.000459 | 0.000518 | -0.000232
390 | 0.000175 | -9.2e-05 | 0.00029 | 0.000308 | -0.000261
390.5 | -0.000342 | 4.8e-05 | 0.000818 | 0.000944 | 0.000391
391 | 3.6e-05 | 0.000661 | 0.000662 | 0.000675 | 0.000197
391.5 | 0.000371 | 0.001107 | 0.000536 | 0.000654 | -7.4e-05
392 | 0.000288 | -0.00049 | 0.000148 | 0.000397 | 0.000355
392.5 | 0.000798 | 4e-06 | 0.000686 | -0.000409 | 0.00015
393 | -7.4e-05 | -7.1e-05 | 0.000644 | 0.000461 | 0.000398
393.5 | -2.3e-05 | 0.000558 | 0.001098 | 0.000472 | 0.000402
394 | -9.5e-05 | 0.000255 | -4e-05 | 0.000176 | 0.000392
394.5 | -7e-06 | 0.000688 | 0.000139 | 0.000875 | -0.000374
395 | -0.000135 | -0.000341 | 0.000716 | 0.000231 | 0.000402
395.5 | -0.000654 | 6.2e-05 | -0.000282 | 0.000839 | 6.1e-05
396 | 0.000222 | -0.000132 | 0.00024 | 0.000619 | 0.000226
396.5 | 0.00047 | -0.000394 | 0.000209 | -0.000137 | 5.6e-05
397 | 0.000514 | -2.4e-05 | 0.000132 | -0.000212 | 0.000339
397.5 | -0.000376 | -0.000439 | -0.000277 | -1.9e-05 | -3.1e-05
398 | -2.5e-05 | 0.0009 | 0.000106 | 7.8e-05 | -0.000171
398.5 | 0.000786 | 0.000194 | -0.00051 | 0.000364 | 0.000249
399 | -3e-06 | -0.000382 | 0.000549 | -0.00016 | -0.000127
399.5 | 0.000175 | -0.000221 | -0.000127 | -0.000346 | -0.000454
400 | -2.9e-05 | 0.000712 | 0.000123 | 0.000922 | -7.5e-05

## _Calculation
Calibration and determination with Excel formulas
Signal = Amplitude at λ — w1 = 335 (measured by Spectro on the processed spectra)
Change nothing here: the formulas show how Spectro turns signals into concentrations. Compare with the _Expected results sheet.
Standard | B (µg/mL) | Signal
Std: A 0, B 4 | 4 | 0.048281
Std: A 0, B 8 | 8 | 0.095999
Std: A 0, B 12 | 12 | 0.14362
Std: A 0, B 16 | 16 | 0.19185
Std: A 0, B 20 | 20 | 0.23988
Slope
Intercept
Correlation coefficient r
Mixture | B taken (µg/mL) | Signal | Found = (Signal − Intercept) / Slope | Recovery %
Mix: A 6, B 10 | 10 | 0.11969
Mix: A 10, B 6 | 6 | 0.071957
Mix: A 14, B 14 | 14 | 0.1679
Mix: A 8, B 16 | 16 | 0.19193
Mix: A 18, B 5 | 5 | 0.060526
Mean recovery %
RSD %

## _Expected results
What Spectro gives for this example — Direct measurement at λmax
Compound: B; template: Direct (zero order, λmax)
Measurement: Amplitude at λ — w1 = 335
Calibration: slope 0.011976, intercept 0.000213, r = 1.00000 (n = 5)
B absorbs alone above 300 nm, so 335 nm (its second band) is free from A.
Mixture | B found | B taken | Rec. %
Mix A 6, B 10 | 9.977 | 10 | 99.77
Mix A 10, B 6 | 5.991 | 6 | 99.84
Mix A 14, B 14 | 14.002 | 14 | 100.01
Mix A 8, B 16 | 16.008 | 16 | 100.05
Mix A 18, B 5 | 5.036 | 5 | 100.72
B: mean recovery 100.08 %, RSD 0.38 %